## Slide 1: Radiation Protection Officer dental radiology - basic
Test exam
Frits Pleiter

## Slide 2
  Question 41
  Equivalent dose rate (in mSv per mAs) at a distance of 1 m from the focus.
  tube voltage thickness aluminum filter
  1 mm 2 mm 3 mm
  50 kV 0.07 0.05 0.03
  70 kV 0.15 0.09 0.05
  90 kV 0.20 0.13 0.08
  Data:
    tube voltage 90 kV
    filter 3 mm aluminum
    tube current 6 mA
    exposure time 0.3 seconds
    distance focus - skin 35 cm
  Calculate the equivalent skin dose per X-ray photo.

## Slide 3
  Evaluation of question 41
  reading figure A at 90 kV and 3 mm aluminum
   0.08 mSv/mAs at 1 meter (1 pt)
  mAs value = 6 mA × 0.3 s = 1.8 mAs (1 pt)
  equivalent dose at 1 meter
  = 0,08 mSv/mAs × 1.8 mAs = 0.14 mSv (1 pt)
  equivalent dose at 35 cm = 0.35 m
  = 0.14 mSv × (1 m / 0.35 m)2 = 1.2 mSv (1 pt)

## Slide 4
  Question 42
  A dental practice has a single intra-oral X-ray device that is subject to registration. For 50% of the photos, the X-ray beam is directed perpendicular to the double-glass window in the outer wall adjacent to the public road.
  Data:
    number of photos per year 2000
    average tube voltage 90 kV
    average mAs value 1,5 mAs per photo
    filter 3 mm aluminum
    distance focus - road 3 meters
    thickness double glass 11 mm
    the contribution of scatter radiation can be neglected
  Check by means of calculation whether or not the shielding effect of the double glass in the window is sufficient to comply with the dose limit for the environment.

## Slide 5
  Evaluation of question 42
  reading figure A at 90 kV and 3 mm aluminum
   0.08 mSv/mAs at 1 meter (1 pt)
  equivalent dose for 2000 / 2 = 1000 photos
   1000 × 0,08 mSv/mAs × 1,5 mAs =
  120 mSv at 1 meter (1 pt)
  equivalent dose at 3 meters
  = 120 mSv × (1 m / 3 m)2 = 13 mSv (1 pt)
  reading figure B at 90 kV and 11 mm glass
   T = 0.3 (1 pt)
  annual effective dose on the street
  = 13 mSv × 0.3 = 4 mSv (1 pt)
  annual limit in case of registration = 10 μSv
   insufficient shielding (1 pt)