## Slide 1: Radiation Protection Officer dental radiography - basic
Risk analysis
Frits Pleiter

## Slide 2
  possible cases
  patient
  unborn child
  dentist
  waiting room
  street
  upstairs neighbours
  accident

## Slide 3
  Patient
  direct beam ?
  equivalent dose ?
  scattering radiation ?
  equivalent dose ?
  effective dose ?
  yes, on the salivary gland(s)
  Hsalivary = Hskin = 5 mSv = 5000 Sv
  yes, on the thyroid
  opening PID = 4 cm  5 cm = 20 cm2
  scattering radiation = Hskin / 1000
  per 100 cm2 at 1 m
  Hscatter = (5000 Sv / 1000)  (20 cm2 / 100 cm2)
  = 5 Sv  0,2 = 1 Sv
  distance skin - thyroid = 0,1 m
  Hthyroid = Hscatter  (1 m / 0,1 m)2
  = 1 Sv  100 = 100 Sv
  wsalivary = 0,01 / 6 = 0,0016 for 1 gland
  wthyroid = 0,04
  Epatient = (0,0016  5000 Sv) + (0,04  100 Sv)
  = 8 Sv + 4 Sv = 12 Sv

## Slide 4
  Unborn child
  direct beam ?
  scattering radiation ?
  equivalent dose ?
  effective dose ?
  no
  yes, on the total body
  scattering radiation according to rule of thumb = 1 Sv per photo at 1 m
  distance skin - child = 0,5 m
  Hchild = 1 Sv  (1 m / 0,5 m)2 = 4 Sv
  wbody = 1
  Echild = 1  4 Sv = 4 Sv
  background = 1600 Sv/y / 365 d/y
  = 4 Sv per day

## Slide 5
  Dentist
  direct beam ?
  scattering radiation ?
  equivalent dose ?
  annual effective dose ?
  no
  yes, on the total body
  scattering radiation according to rule of thumb = 1 Sv per photo at 1 m
  distance skin - dentist = 3 m
  no shielding
  Hdentist = 1 Sv  (1 m / 3 m)2
  = 0,1 Sv
  wbody = 1
  number of photos = 2000 per year
  Edentist = 1  0,1 Sv  2000
  = 200 Sv
  annual limit = 1 mSv

## Slide 6
  Waiting room
  direct beam ?
  equivalent dose ?
  scattering radiation ?
  annual effective dose ?
  perhaps
  Hskin = 5 mSv
  distance focus - skin = 0,2 m
  distance focus - waiting room = 2 m
  no shielding
  Hwaiting room = 5 mSv  (0,2 m / 2 m)2
  = 0,05 mSv
  yes, but the contribution is negligible
  wbody = 1
  number of photos = 2000 per year, half of which in the direction of the waiting room
  Ewaiting room = 1  0,05 mSv  (2000 / 2)
  = 50 mSv
  annual limit = 1 mSv

## Slide 7
  Street
  direct beam ?
  equivalent dose ?
  scattering radiation ?
  annual effective dose ?
  perhaps
  Hskin = 5 mSv
  distance focus - skin = 0,2 m
  distance focus - waiting room = 3 m
  no shielding
  Hstreet = 5 mSv  (0,2 m / 3 m)2
  = 0,02 mSv
  yes, but the contribution is negligible
  wbody = 1
  number of photos = 2000 per year, half of which in the direction of the street
  Estreet = 1  0,02 mSv  (2000 / 2)
  = 20 mSv
  annual limit = 10 Sv

## Slide 8
  Upstairs neighbours
  direct beam ?
  scattering radiation ?
  equivalent dose ?
  annual effective dose ?
  no
  yes, on the total body
  scattering radiation according to rule of thumb = 1 Sv per photo at 1 m
  distance skin - neighbours = 3 m
  Hneighbours = 1 Sv  (1 m / 3 m)2
  = 0,1 Sv
  wbody = 1
  number of photos = 2000 per year
  Eneighbours = 1  0,1 Sv  2000
  = 200 Sv
  annual limit = 10 Sv

## Slide 9
  Accident
  unintended unlikely
  not impossible
  scenario
  thyroid and all
  six salivary glands
  in the direct beam
  equivalent dose ?
  effective dose?
  Hskin = 5 mSv
  Hsalivary = Hthyroid = Hskin
  wsalivary = 0,01
  wthyroid = 0,04
  Eaccident = (0,01 + 0,04)  5 mSv = 0,25 mSv
  annual limit = 1 mSv
  thus no reason to panic